## Slide 1
PHASE NOISE MEASUREMENTS
Dave Robinson, WW2R
Microwave Update 2004

## Slide 2: MEASUREMENTS MADE
10MHz 5
11MHz-1GHz 12
1GHz+ 9

## Slide 3: LESS THAN 1GHz
VERY SIMILAR RESULTS: MEASURING ANALYSERS CHARACTERISTICS, MASKING OSCILLATOR PHASE NOISE

## Slide 4: AFFECT OF CRYSTAL DRIVE LEVEL

## Slide 5: 5616MHz Brick locked to GPS with Shera GPS lock box reference

## Slide 6: 5616MHz Brick DMK locked to GPS via Z3801A

## Slide 7: 5616 MHz Brick (DMK disabled)

## Slide 8: FREQUENCY WEST BRICK 10224MHz

## Slide 9: “Qualcom Texas Synth” multiplied to 10224MHz
Note higher reference level (-30dBc/Hz)

## Slide 10: JWM 5112Mhz SOURCE

## Slide 11: JWM 1152MHz + x9 mult
Problem with unit? Produced signal around 850MHz on one occasion.

## Slide 12: DEMI prototype (xtal osc+mult)

## Slide 13: QUALCOM 1152MHz SYNTH

## Slide 14: What is Phase Noise?
Phase noise is ‘random’ noise energy near an RF carrier such as a local oscillator signal. Phase noise is measured in dBc/Hz at a given frequency offset.
As the carrier frequency is increased through the LO chain, the phase noise is scaled by:
20Log(n), where n = ratio of frequencies
For example, a 10MHz signal with phase noise of -105dBc/Hz @ 1KHz would be ‘the same’ as a 1150MHz signal with -63.8dBc/Hz @ 1KHz phase noise if measured at 1150MHz.
For amateur radio purposes, phase noise can be classified as either “Near” or “Far”.
(From wa1zms presentation)

## Slide 15: KE5FX SYNTH 1136MHz
-91.83 dBC/Hz at 10kHz

## Slide 16: KE5FX SYNTH + DOUBLER
-91.83+ 20 LOG (2) = -91.83+6= -85.83DBC/HZ

## Slide 17: KE5FX SYNTH + DOUBLER
-91.83+ 20 LOG (2) = -91.83+6= -85.83DBC/HZ

## Slide 18: KE5FX+ X9 MULTIPLIER
-91.83+ 20 LOG (9) = -91.83+19.08= -72.8dBC/HZ

## Slide 19: KE5FX+ X9 MULTIPLIER
-91.83+ 20 LOG (9) = -91.83+19.08= -72.8dBC/HZ

## Slide 20: Future work
Lots more oscillators to be measured as time allows.
Better measurement techniques needed to measure phase noise below 1GHz